## Sheet1
HPAFM Wiring
Code | Meaning
C | Connector (specific plugs)
P | Processor device
R | Relay
H | Heater
T | Transformer
L | Light
S | Pushbutton Switch
M | Measurement Device
Power Wiring
Name | Description | Interface | Rating | Connects to | Part Number
T1 | Isolating UPS | NEMA 125V/15A plugs | 1000VA 700W |  | Tripp Lite OMNI1000ISO
S1 | Emergency off key switch | Bare wire | 120VAC/6A (x2) |  | McMaster 6785K26
Sensors Wiring
Name | Description | Interface | Rating | Connects to | Part Number
PresG | Pressure gauge connector | 4-pin, blue wire = +12V, brown wire = 0V
LEDs | LED connector | 4-pin, blue=+12V, Brown=Red-, White=Green-, Black=Blue-
HtrBox | Heater box connector | 4-pin, blue=relay on+, brown=relay on-, black and white=current transformer
T2 | Current transformer 1/20 | 2-wire AC output | 0-5A out |  | McMaster 65795K82
L1 | LED strip 3 color | 4-wire (anode or cathode common?)
R1 | Heater relay | Control: DC 4 to 32V, output: AC 24 to 280V | 280VAC, 10A out |  | Omega SSRDIN280DC10
T3 | Heater 240VAC transformer | NEMA 125V/15A plug, british plug | 500W in, fused |  | Ted Pella 54585
M1 | Pressure gauge 4-20mA 3000 PSI | 4-20mA, 8-32VDC loop power | 8-32VDC, reverse protected |  | Omega DPG1000L-3KG
M2 | Thermocouple for water temp | Thermocouple K-type | Connector 220C |  | Omega KMQIN-062G-6
M3 | Thermocouple for surface temp | Thermocouple K-type | Connector 220C |  | Omega KMQIN-062G-6
C1 | Thermocouple extension cable | Thermocouple K-type | Connector 220C |  | Omega RECK1-9
C2 | Thermocouple extension cable | Thermocouple K-type | Connector 220C |  | Omega RECK1-9
H1 | Heater top | 2-wire heater wire | 240VAC, 100W |  | Omega MSH00029
H2 | Heater bottom | 2-wire heater wire | 240VAC, 100W |  | Omega MSH00029
C3 | Heater top hot wire | Heater wire | 450C, 300V, 16ga |  | Omega HTMG-1CU-316S/C
C4 | Heater bottom hot wire | Heater wire | 450C, 300V, 16ga |  | Omega HTMG-1CU-316S/C